Lumb and Jones - Chap 17 p474-492

Maintenance of vaporizers

- Follow manufactures instructions regarding care of vaporizers
- Servicing includes: evaluation of operation, cleaning, changing filters, replacing worn parts, recalibration.
- Halothane and methoxyflurane contain preservatives that can collect and affect function of the vaporizer

Use of the wrond anesthetic in a an agent-specific vaporizers

- Low output of anesthetic is exoected with a lower vapor pressure anesthetic into a higher vapor pressure vaporizer
- High and potentially lethal output possible is anesthetic with higher vapor pressure is placed into machine designed for lower vapor pressure.
- Halothane and isoflurane are similar in vapor pressure so they are reasonable the same, however the machine should remain agent specific, and at the very least the halothane machine should be recalibrated for the use of isoflurane (or vice versa).

Common gas outlet

- the outlet is the site from which gases that have passed thru the flowmeter, vaporizer (VOC), and fludh valve exit the anesthesia machine on the way to the breathing system
- typically there are 15mm internal diameter opening to which a rubber tube attaches
- the other end of the tubing circuit connects to the fresh gas inlet of the breathing circuit.
- Disconnecting the coon gas outlet, the vaporizer outlet, or the fresh gas outlet will result in loss of gas flow to the breathing circuit. The disconnection of the first two may not be immediately recognized during use
  - with circle system: low FIO2, increased respiratory effort, light anesthesia likely.
  - Non-rebreathing system: rebreathing of exhaled CO2 will occur, low FIO2 likely, light anesthesia and increased RE expected.

Breathing systems

- breathing systems add resistance to the patients inhalation/exhalation
- diameter of tubes and other conduits are major factor in determining the resistance
  - doubling the radius decreases the resistance 16x.
  - Halving the length halves the resistance
  - changing the direction of flow of gas thru restrictive orifices creates turbuent flow and increases resistance
- tubes should be as short as practical, max diametes in the conduits, fewest bends and restrictions in the path of gas flow.

Systems using chemical absorption of CO2 –

rebreathing systems - Part or all of the exhaled gases (after removal of CO2), flow back to the patient.
- Circle
  - pediatric - <6.8kg
  - standard adult 6.8kg – 135kg
  - Large animal >135kg
  - gases move only in one direction – exhaled gases enter the Y piece and flow thru the expiratory breathing tube and the expiratory one-way valve. Gases enter the reservoir bag before or after coursing thru the CO2 canister. On inspiration, gases exit the reservoir bag and travel thru the inspiratory one-way valve, the inspiratory breathing tube and the y-piece to the patient.
  - Components:
    - y piece: contributes to dead space. Attached to the ET tube.
    - Breathing tubes: should have internal diameter larger than the ET tube. Does not contribute to mechanical dead space if the one way valves are functional
    - one way valves: paired, direct flow away from the patient on expiration and towards the patient on inhalation
    - fresh gas inlet: where gases enter the system (from the common gs outlet or vaporizer outlet)
    - pop off valve: vents gases to the scavenger system to prevent buildup of excessive pressure within the circle. Allows rapid anesthetic gas elimination. Should vent 1-2 cm H2O when fully opened.
    - Reservoir bag: located on the absorber side of the circle. Provides way to provide mechanical or hand ventilation. Allows assessment of spontaneous ventilation and roughly estimate tidal volume. Should not allow pressures to exceed 60cmH2O, so provides immediate relief for a closed pop off valve. Should be 6x the patients tidal volume.
    - Manometer: pressure gauge attached to the top of the absorber assembly. In cmH2O. Used to assess pressures during assisted or controlled ventilation
    - air intake valve: allows room ait to enter the circle and prevents a patient from inspiring against a negative pressure and becoming hypoxic.
    - In-circuit vaporizers: patients must inhale thru the vaporizers
    - absorber: Canister for CO2 absorbing. Located between the one way valves, exhalation side.
- to-and-fro
Chemical absorption of CO2

- Calcium hydroxide – primary component of soda lime and barium hydroxide lime
- small granules offer most surface are for chemical reactions, larger granules impose less resistance to gas flow thru the canister.
- Unused granules are soft, expended ones hard (conversion to calcium carbonate)
- pH indicators added – color change to violet to indicate granules are exhausted. (may revert to white during storage).
- 2 NaOH + 2 H2CO3 + Ca (OH)2 → CaCO3 + NaCO3 + 4 H2) + heat
Fresh gas flows for circle breathing systems

- closed circle system – low flow of anesthetic gases and O2 to the patient and system. O2 approximates the oxygen consumption. ~4-11ml/kg/min
  - N2O generally not used
- Low-flow circle – greater O2 flow rates than patient consumption. 4-7 ml/kg/min. Disadvantage: mask induction, may need higher flow rates initially.
- Semi-closed circle system – fresh gas exceeds the uptake of O2 and anesthetic by patient. 22-44 ml/kg/min.

To-and-fro systems

- CO2 canister between ET tube and reservoir bag. Pop off valve and fresh gas inlet between the canister and ET tube connector.
- Lower flow rates – CO2 eliminated by chemical absorption
- higher flow rates – CO2 elimination thru pop off valve.
- Heat produced during CO2 absorption may be transferred to the patient during inspiration.

Mapleson system

- No chemical absorbent for CO2, depend on high fresh gas rates
- “non-rebreathing systems”
- More expensive, contribute to hypothermia
- Magill, Bain, Ayer's T-peice, Norman mask elbow

Magill system

- “Mapelson A”
- fresh gas inlet, overflow valve near patient, corrugated tube connecting the patient nd to a reservoir bag.
- Some rebreathing of CO2 occurs during assisted ventilation

Bain coaxial system

- “Mapelson D”
- tube within a tube. Internal tube delivers fresh gases to patient, external tube removes waste gases. Reservoir bag attached to the external corregated tube.
- 200-300 ml/kg/min human, 100-150ml/kg/min for veterinary patients. Others suggest 200ml/kg/min for <7kg. ~1.5 to 3 x minute volume.
- Risk of rebreathing CO2 if insufficient gas rates are used.

Ayre's T piece and Norman mask elbow systems

“Mapelson F” with reservoir bag, “E” without

Systems with non-rebreathing valves

- cumbersome, seldom used in veterinary medicine
- Stephens-Slater system
- resuscitation bags – Ambu bag

closed containers and masks

- Chamber: Glass, plexiglass induction chambers, airtight. High flow rates into chamber, gas outlet.
- Masks: should fit snugly over the muzzle, attached to breathing system to provide reservoir of gases to meet peak inspiratory flow demands. Well ventilated areas.

Scavenging waste anesthetic gases
- AVMA liability insurance trust: Numerous studies in US and abroad have found no conclusive evidence that waste gases of trace amounts of waste gases cause specific health problems. Evidence to suggest that removal of gases from veterinary facilities will likely improve the occupational health of veterinary staff
- OSHA – no limits on exposure to anesthetics
- NIOSH – 2 ppm for halogenated hydrocarbon (isoflurane, halothane)
- American conference of governmental industrial hygienists – 50ppm for halothane and N2O as an 8 hour time-weighted average.

Recommendations for controlling waste gases:

- Personnel education
- scavenger systems
- ventilation
- maintenance of machines
- log usage, maintenance, performance
- minimize spillage
- Ways to decrease contamination of the occupational environment with anesthetic gases:
  - Avoid spills when filling
  - start gas after patient intubation
  - use inflated cuff ET tubes
  - Occlude Y piece when disconnecting
  - use scavenging pop off valve
  - discharge all gases thru pop off
  - flush breathing system with O2 before disconnecting
  - use minimum gas flows
  - minimize use of containers and masks for induction
  - proper ventilation

Scavenging systems

- Passive or active
  - passive – pipe directly into the atmosphere and absorption devices
    - canisters – can be expensive, need to be changed regularly, do not work with N2O.
  - active – piped vacuum and active duct systems
    - central vacuum systems

Anesthesia Checklist – daily

- Check central O2 and N2O supplies
- inspect flowmeters, vaporizers, gauges, supply hoses.
- Assure that the waste scavenging system is properly attached
- Assure vaporizer is properly filled
- Check O2 cylinders for contents and leaks
- Check N2O supply
- Test flowmeters
- Test central pipeline supplies of O2 and N2O
- Check valves in a circle system
- Leak test machine
- Test pop off valve

Lumb and Jones - Chap 17 p474-492

Questions:

What is the expected result of using the incorrect anesthetic in a vaporizer?

What is the effect of halving the length of tubing in an anesthetic breathing system?

what is the difference between rebreathing and nonrebreathing systems? Name an example of each. Name an advantage of each.

Label the diagram:

A: _____________
B: inspiratory / expiratory side (choose one)
C: Inspiratory / expiratory side (choose one)
D: ______________

Name the breathing system depicted in the image:

____________
B

D

A

C

What is an appropriate oxygen flow level for a Bain system? What is a consequence of too low a flow level?

Lumb and Jones - Chap 17 p474-492

Questions:

1. What is the expected result of using the incorrect anesthetic in a vaporizer?
  Low output of anesthetic – if an anesthetic with a lower vapor pressure is placed into a vaporizer designed for a drug with a higher vapor pressure.
  Higher, potentially lethal amounts of anesthetic can be released if a higher vapor pressure anesthetic is placed in a vaporizer designed for a lower vapor pressure anesthetic.

1. What is the effect of halving the length of tubing in an anesthetic breathing system?
  Halves the resistance to breathing

1. what is the difference between rebreathing and nonrebreathing systems? Name an example of each. Name an advantage of each.
  Rebreathing: All or part of the exhaled gases are returned to the patient after CO2 extraction.
    Example: Circle and to-and-fro
    Advantage: anesthetic conservation, more expensive to purchase but less expensive to operate.
  Non-rebreathing: Exhaled gases are not returned to the patient after CO2 extraction
    Example: Stephens-Slater, Fink, Digby-Leigh, closed chamber, mask, (resuscitation bag)
    Advantage: no rebreathing of exhaled gases

Label the diagram:

A: __Y-piece___________
B: inspiratory / expiratory side (choose one)
C: Inspiratory / expiratory side (choose one)
D: __CO2 absorber____________

Name the breathing system depicted in the image:

Circle

1. What is an appropriate oxygen flow level for a Bain system? What is a consequence of too low a flow level?
  200-300 ml/kg/min human, 100-150ml/kg/min for veterinary patients. Others suggest 200ml/kg/min for <7kg. ~1.5 to 3 x minute volume.
  Rebreathing of CO2.